## Sheet1
|  | TANZANIA BUREAU OF STANDARDS |  |  |  |  | CONTROLLED
METROLOGY LABORATORY |  | METROLOGY LABORATORY
Document MET-TF-03 |  | Tittle: Calibration of Timer |  | Revision: 0 |  | Page: 1 of 1 |  | Copy number:
|  |  |  | Issue: 1 |  |  |  | Issued to:
Document type: Uncertainty Budget |  | Prepared by: Authorized personnel in time & frequency laboratory. |  | Approved by: HML |  | Effective date: 2020-01-31 |  | Issued date: 2020-02-07 | 2020-02-07
|  | Signature: |  | Signature:
UNCERTAINTY BUDGET FOR TIMER
Customer's name: | Lordrick Apps
Equipment: | TIMER
ID no: | TBS/LAB/01
Certificate No: | 34552
Job No: | 2021 - 999
Calibrated by: | J. Yarrot
|  |  | Time interval (s) |  | <1800s
Source of uncertainity | Expected value | Distribution | Uncertainty(s) | Divisor | u(xi) (s) | Ci | u(yi) (s) | Reliability | νi
Human reaction time standard dev. | 0 | Normal | 0.0125 | 1 | 0.0125 | 1 | 0.0125 | 95 | 200
Human reaction Time bias | 0 | Normal | 0.013 | 1 | 0.013 | 1 | 0.013 | 95 | 200
Standard stopwatch | 0 | Normal | 0.04 | 2 | 0.02 | 1 | 0.02 | 95 | 20000%
Resolution of UUT | 0 | Rectangular | 0.5 | 1.7321 | 0.28868 | 1 | 0.28868 | 95 | 200
Drift of STD | 0 | Rectangular | 0.004 | 1.7321 | 0.0023094 | 1 | 0.0023094 | 95 | 200
[5 empty] | Level of Confidence: |  | 95%
[5 empty] | Combined uncertainty |  | 0.28994 | Veff | 203.5159
[5 empty] | Expanded uncertainty |  | 0.58347 | Student-t | 2.0124
U= ± 1
Reaction time No. of measurements |  | 3